Ventilator Modes Outline

Overview:

- Our goal with various ventilator modes is to deliver an adequate minute ventilation to the patient while minimizing any ventilator associated lung injuries that can occur.

- Minute Ventilation = Respiratory Rate x Tidal Volume.

- If different modes, we can manipulate various components such as the tidal volume, respiratory rate, inspiratory to expiratory ratio, positive end expiratory pressure (PEEP), flow, and several other characteristics.

Volume vs. Pressure Modes:

- Volume controlled ventilation is the idea that the ventilator will push a preset volume into the patient and will give this volume regardless of the pressure inside the lungs.

  - Compliance is the idea that a container will be able to accept more volume with a limited change in pressure.

  - Think of the lungs as a balloon that you are filling up. The more compliant the balloon, the easier it is to put volume into.

  - However, if the balloon already has a high pressure or is less compliant, then it will be difficult to put the same volume of air into the balloon.

  - If the lungs have a normal compliance, then the ventilator will be able to push in the preset volume with limited change in pressure.

  - However, if the lungs are less compliant, the ventilator will still try and push the volume of gas into the lungs and cause the pressure to skyrocket.

  - During volume control ventilation, the inspiratory flow going through the ventilator will be the same throughout the inspiratory cycle.

- Pressure control ventilation delivers a preset inspiratory pressure that is given over a selected time interval, meaning, the volume delivered will change based upon when the ventilator registers a certain pressure has been reached.

  - If the compliance of the lungs are low or the airway resistance is high, the pressure limit will be reached quickly during inspiration and only a small tidal volume breath will be given to the patient.

  - However, if compliance is good and resistance low, the patient can get a massive tidal volume before the pressure is reached (watch out so that they don’t get too big of breaths!!)

  - Since the volume it takes for the ventilator to reach the pressure limit will vary with each breath based on, the ventilator changes the inspiratory flow with each breath to try and make the patient gets the breath over the same inspiratory time with each breath.

  - Pressure control ventilation is useful for patients that are having high airway pressures. This mode allows the inspiratory flow to change from breath to breath and maximizes the amount of volume per the max airway pressure. This also decreases the risk of barotrauma to the patient.

  - Another advantage is the ability to compensate for leaks in the patient- ex: if you don’t have a cuff on the ETT.

  - Disadvantages of pressure-controlled ventilation include lower tidal volumes delivered to the patient so they can start to become hypoxic if you don’t supplement with higher respiratory rate to maintain the minute ventilation.

  - Additionally, you have to keep a closer eye on the patient after changing positions or doing anything surgically that would alter the compliance of the lungs.

    - The airway pressure could increase if the ETT becomes kinked which would cause the pressure to be reached sooner and a lower tidal volume given to the patient.

    - If the patient is put into Trendelenburg position (head down), the abdominal contents compress on the lungs and cause a pressure limit to be reached with a lower volume. This is also true during laparoscopic cases when there is a pneumoperitoneum used to inflate the abdomen.

    - These cases would cause for a low tidal volume and require a higher RR to maintain minute ventilation.

    - On the flip side, when coming out of these high-pressure scenarios (releasing the pneumoperitoneum, giving a bronchodilator, leveling back the patient’s bed, etc. can cause the tidal volume to skyrocket due to less pressure.

    - You have to be mindful during this point of a case to watch how high the tidal volumes are.  You can lower the pressure limit to reduce the volume given.

Ventilator sub-modes:

- Pressure Control Ventilation with Volume Guarantee (PCV-VG) -

  - Combines the benefits of both pressure control and volume control ventilation.

  - You set a minimum volume you want to patient to receive along with the RR. The ventilator will adapt with each breath to give the volume you set by using the least amount of pressure to get the given volume.  It manipulates the inspiratory flow and pattern to limit the pressure.

  - This means you don’t have to worry about barotrauma as much as with simple volume control ventilation and at the same time you don’t have to worry about their lungs getting drastically different volumes when the patient position is changed.

- Controlled Mandatory Ventilation (CMV) -

  - Delivers a preset breath at whatever length you set and whatever rate you set regardless of if the patient is trying to breathe in between breaths.

  - This mode is best used if the patient is paralyzed or apneic, so they are not bucking/fighting the vent.

- Assist Control ventilation (AC) -

  - Same as controlled mandatory ventilation but if the patient starts to breath in between breaths, the ventilator will not fight the patient but rather give the full predetermined volume.

  - So, if the patient is breathing over the respiratory rate you have set, they might hyperventilate due to their minute ventilation being increased.

- Synchronized Intermittent Mandatory Ventilation (SIMV) -

  - Same as the two previous modes but when the patient tries to take a breath on their own, the ventilator will adjust its next breath to coordinate better with the patient’s breathing pattern and the patient’s breath will not get the full preset volume but rather a supplement breath with PS

  - AKA the more breaths the patient takes, the ventilator will let the patient breath on their own, as long as they are reaching the minimum required minute ventilation.

- Pressure Support Ventilation PRO (PSV-PRO) -

  - Pressure support simply helps each patient spontaneous breathe with a preset pressure support. The patient has to initiate each breath.

  - You can set how much support pressure the patient will receive to augment each breath along with the trigger.

  - The trigger is how much work the patient has to do before the ventilator will pick it up as a patient breath. The lower the trigger, the less amount of work the patient has to do before the ventilator will deliver the pressure support.

  - If you use a PRO mode, then the ventilator will watch for apnea and will revert to an SIMV mode by supplying a back up RR to the patient if the patient becomes apneic.

- CPAP vs BiPAP -

  - Continuous positive airway pressure (CPAP) gives a continual pressure throughout the entire respiratory breathing pattern, both inspiration and expiration. This is used during inspiration like a PS mode but maintains the pressure during expiration to keep the airway and alveoli open.

  - Biphasic positive airway pressure (BiPAP) provides the inspiratory pressure support during a spontaneous breath and then a different pressure during expiration to keep the airway/alveoli open. Similar to CPAP, but there are two different pressures (one during inspiration and one during expiration).

Inspiratory to Expiratory Ratio I:E -

- A normal I:E ratio is 1:2 which means there is twice as much time spent in expiration than inspiration.

- You can manipulate this; for example if you have a patient with severe asthma, you can change the ratio to 1:3 which will give you longer expiration times and give the patient more time to exhale the gas.

- Inverse Ratio Ventilation (IRV) switches the I:E to make inspiration longer than expiration. This is used in patients with ARDS and requires the patient to be completed paralyzed and sedated.

- This allows for lower peak airway pressures and improved oxygenation but will decrease the venous return and cardiac output.

- Can be at risk for building up CO2 due to less expiration time and cause auto-peep/breath stacking to occur by not allowing enough time for the gas to leave the lungs before the next breath comes in.